|  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  | ***CALIBRATION CERTIFICATE NO: PPE078SW007201/001*** | | | | | | |
|  |  |  |  |  |  |  |  |  |  |  |  |
| **CLIENT NAME :** | | | **Rusan Pharma Limited** | | | | | | | |  |
| **ADDRESS:** | |  | Khasra No.: 122 MI, Central Hope Town, Behind Pharma City, Selaqui, Dehradun-248197, INDIA | | | | | | | | |
|  |  |  |
| **ENVIRONMENTAL CONDITION:** | | | **AMBIENT TEMPERATURE & HUMIDITY** | | | | |  |  |  |  |
| ***INSTRUMENT DETAILS*** | | | | | | | | | | | |
| **INSTRUMENT** | | Stop Watch | | | | **DEPARTMENT** | | Quality Control | | | |
| **MAKE** | | Racer | | | | **CAL. RANGFE** | | 2 Hours | | | |
| **RANGE** | | 0 – 23 Min : 59 Sec | | | | **LOCATION** | | Q.C. Lab | | | |
| **ACCURACY** | | ± 1% of Full Scale | | | | **TAG ID** | | SW/QCE/001 | | | |
| **LEAST COUNT** | | 1 Sec. | | | | **CAL. DATE** | | 09/07/2021 | | | |
| **UNIT OF MEASURMENT** | | Sec. | | |  | **DUE ON** | | 08/07/2022 | | | |
| **CALIBRATION PROCEDURE NUMBER** | | | | | | PPE/SOP/TIME/TM/01-00 | | | | | |
| **OBSERVATION TABLE** | | | | | | | | | | | |
| **Sr. No** | **STANDARD READINGS ON UNIT UNDER CALIBRATION (HH:MM:SS: 1/100s)** | | | **OBSERVED READINGS ON STANDARD INSTRUMENT (HH:MM:SS:1/100s)** | | | | **ERROR (HH:MM:SS:1/100s)** | | | |
|
|
| **1** | 0:00:00:00 | | | 0:00:00:02 | | | | 0:00:00:02 | | | |
| **2** | 0:30:00:00 | | | 0:30:00:04 | | | | 0:00:00:04 | | | |
| **3** | 1:00:00 | | | 1:00:00 | | | | 0:00:00 | | | |
| **4** | 2:00:00 | | | 2:00:00 | | | | 0:00:00 | | | |
| **Note: (1/100 s  reading is available till 1 Hours)** | | | | | | | | | | | |
| **REMARK: THE STOP WATCH IS CALIBRATED SUCCESSFULLY WITHIN THE SPECIFIED LIMITS** | | | | | | | | | | | |
|  |  |  |  |  |  |  |  |  |  |  |  |
| ***REFERENCE STANDARDS USED*** | | | | | | | | | | | |
| **INSTRUMENT** | | Digital Stop Watch | | | | **RANGE** | | 0-24 Hour | | | |
| **MAKE** | | Racer | | | | **REPORT NO** | | 2021/01/0061 | |  |  |
| **INST ID. NO** | | PPE/C/SW/01 | | | |  |  | **(Traceable To NABL)** | |  | |
|  |  |  |  |  |  | **DUE DATE** | | 06/01/2022 | |  |  |
|  | | | |  | | | |  | | | |
|
|
|
| **CALIBRATED BY** | | | | **CHECKED & APPROVED BY** | | | | **REVIEWED BY** | | | |
| ***NOTE*** | |  |  |  |  |  |  |  |  |  |  |
| This report refers only to the particular item calibrated at site/laboratary. | | | | | | | | | | |  |
| Pin Point Engineers is not liable for any change in calibration data & performance specifications on account of malfunctioning the standards/instrument/equipment covered by this report due to damage caused to it after issuance of this report. | | | | | | | | | | | |
|
| The Standard(s) used for calibration is/are calibrated by using reference standards traceable to National/International standards. | | | | | | | | | | | |
| The calibration results reported are valid at the time of and under the conditions of measurement. | | | | | | | | | | | |
| This report may not be reproduced, except in full, without the prior permission of "Pin Point Engineers". | | | | | | | | | | | |
| Any correction in this certificate invalidated the certificate. | | | | | | | | | | | |
|  |  |  |  |  | ***CALIBRATION CERTIFICATE NO: PPE078SW007201/002*** | | | | | | |
|  |  |  |  |  |  |  |  |  |  |  |  |
| **CLIENT NAME :** | | | **Rusan Pharma Limited** | | | | | | | |  |
| **ADDRESS:** | |  | Khasra No.: 122 MI, Central Hope Town, Behind Pharma City, Selaqui, Dehradun-248197, INDIA | | | | | | | | |
|  |  |  |
| **ENVIRONMENTAL CONDITION:** | | | **AMBIENT TEMPERATURE & HUMIDITY** | | | | |  |  |  |  |
| ***INSTRUMENT DETAILS*** | | | | | | | | | | | |
| **INSTRUMENT** | | Stop Watch | | | | **DEPARTMENT** | | Quality Control | | | |
| **MAKE** | | Racer | | | | **CAL. RANGFE** | | 2 Hours | | | |
| **RANGE** | | 0 – 23 Min : 59 Sec | | | | **LOCATION** | | Q.C. Lab | | | |
| **ACCURACY** | | ± 1% of Full Scale | | | | **TAG ID** | | SW/QCE/002 | | | |
| **LEAST COUNT** | | 1 Sec. | | | | **CAL. DATE** | | 09/07/2021 | | | |
| **UNIT OF MEASURMENT** | | Sec. | | |  | **DUE ON** | | 08/07/2022 | | | |
| **CALIBRATION PROCEDURE NUMBER** | | | | | | PPE/SOP/TIME/TM/01-00 | | | | | |
| **OBSERVATION TABLE** | | | | | | | | | | | |
| **Sr. No** | **STANDARD READINGS ON UNIT UNDER CALIBRATION (HH:MM:SS: 1/100s)** | | | **OBSERVED READINGS ON STANDARD INSTRUMENT (HH:MM:SS:1/100s)** | | | | **ERROR (HH:MM:SS:1/100s)** | | | |
|
|
| **1** | 0:00:00:00 | | | 0:00:00:01 | | | | 0:00:00:01 | | | |
| **2** | 0:30:00:00 | | | 0:30:00:03 | | | | 0:00:00:03 | | | |
| **3** | 1:00:00 | | | 1:00:00 | | | | 0:00:00 | | | |
| **4** | 2:00:00 | | | 2:00:00 | | | | 0:00:00 | | | |
| **Note: (1/100 s  reading is available till 1 Hours)** | | | | | | | | | | | |
| **REMARK: THE STOP WATCH IS CALIBRATED SUCCESSFULLY WITHIN THE SPECIFIED LIMITS** | | | | | | | | | | | |
|  |  |  |  |  |  |  |  |  |  |  |  |
| ***REFERENCE STANDARDS USED*** | | | | | | | | | | | |
| **INSTRUMENT** | | Digital Stop Watch | | | | **RANGE** | | 0-24 Hour | | | |
| **MAKE** | | Racer | | | | **REPORT NO** | | 2021/01/0061 | |  |  |
| **INST ID. NO** | | PPE/C/SW/01 | | | |  |  | **(Traceable To NABL)** | |  | |
|  |  |  |  |  |  | **DUE DATE** | | 06/01/2022 | |  |  |
|  | | | |  | | | |  | | | |
|
|
|
| **CALIBRATED BY** | | | | **CHECKED & APPROVED BY** | | | | **REVIEWED BY** | | | |
| ***NOTE*** | |  |  |  |  |  |  |  |  |  |  |
| This report refers only to the particular item calibrated at site/laboratary. | | | | | | | | | | |  |
| Pin Point Engineers is not liable for any change in calibration data & performance specifications on account of malfunctioning the standards/instrument/equipment covered by this report due to damage caused to it after issuance of this report. | | | | | | | | | | | |
|
| The Standard(s) used for calibration is/are calibrated by using reference standards traceable to National/International standards. | | | | | | | | | | | |
| The calibration results reported are valid at the time of and under the conditions of measurement. | | | | | | | | | | | |
| This report may not be reproduced, except in full, without the prior permission of "Pin Point Engineers". | | | | | | | | | | | |
| Any correction in this certificate invalidated the certificate. | | | | | | | | | | | |
|  |  |  |  |  | ***CALIBRATION CERTIFICATE NO: PPE078SW007201/001*** | | | | | | |
|  |  |  |  |  |  |  |  |  |  |  |  |
| **CLIENT NAME :** | | | **Rusan Pharma Limited** | | | | | | | |  |
| **ADDRESS:** | |  | Khasra No.: 122 MI, Central Hope Town, Behind Pharma City, Selaqui, Dehradun-248197, INDIA | | | | | | | | |
|  |  |  |
| **ENVIRONMENTAL CONDITION:** | | | **AMBIENT TEMPERATURE & HUMIDITY** | | | | |  |  |  |  |
| ***INSTRUMENT DETAILS*** | | | | | | | | | | | |
| **INSTRUMENT** | | Stop Watch | | | | **DEPARTMENT** | | Quality Control | | | |
| **MAKE** | | Racer | | | | **CAL. RANGFE** | | 2 Hours | | | |
| **RANGE** | | 0 – 23 Min : 59 Sec | | | | **LOCATION** | | Q.C. Lab | | | |
| **ACCURACY** | | ± 1% of Full Scale | | | | **TAG ID** | | SW/QCE/001 | | | |
| **LEAST COUNT** | | 1 Sec. | | | | **CAL. DATE** | | 09/07/2021 | | | |
| **UNIT OF MEASURMENT** | | Sec. | | |  | **DUE ON** | | 08/07/2022 | | | |
| **CALIBRATION PROCEDURE NUMBER** | | | | | | PPE/SOP/TIME/TM/01-00 | | | | | |
| **OBSERVATION TABLE** | | | | | | | | | | | |
| **Sr. No** | **STANDARD READINGS ON UNIT UNDER CALIBRATION (HH:MM:SS: 1/100s)** | | | **OBSERVED READINGS ON STANDARD INSTRUMENT (HH:MM:SS:1/100s)** | | | | **ERROR (HH:MM:SS:1/100s)** | | | |
|
|
| **1** | 0:00:00:00 | | | 0:00:00:02 | | | | 0:00:00:02 | | | |
| **2** | 0:30:00:00 | | | 0:30:00:04 | | | | 0:00:00:04 | | | |
| **3** | 1:00:00 | | | 1:00:00 | | | | 0:00:00 | | | |
| **4** | 2:00:00 | | | 2:00:00 | | | | 0:00:00 | | | |
| **Note: (1/100 s  reading is available till 1 Hours)** | | | | | | | | | | | |
| **REMARK: THE STOP WATCH IS CALIBRATED SUCCESSFULLY WITHIN THE SPECIFIED LIMITS** | | | | | | | | | | | |
|  |  |  |  |  |  |  |  |  |  |  |  |
| ***REFERENCE STANDARDS USED*** | | | | | | | | | | | |
| **INSTRUMENT** | | Digital Stop Watch | | | | **RANGE** | | 0-24 Hour | | | |
| **MAKE** | | Racer | | | | **REPORT NO** | | 2021/01/0061 | |  |  |
| **INST ID. NO** | | PPE/C/SW/01 | | | |  |  | **(Traceable To NABL)** | |  | |
|  |  |  |  |  |  | **DUE DATE** | | 06/01/2022 | |  |  |
|  | | | |  | | | |  | | | |
|
|
|
| **CALIBRATED BY** | | | | **CHECKED & APPROVED BY** | | | | **REVIEWED BY** | | | |
| ***NOTE*** | |  |  |  |  |  |  |  |  |  |  |
| This report refers only to the particular item calibrated at site/laboratary. | | | | | | | | | | |  |
| Pin Point Engineers is not liable for any change in calibration data & performance specifications on account of malfunctioning the standards/instrument/equipment covered by this report due to damage caused to it after issuance of this report. | | | | | | | | | | | |
|
| The Standard(s) used for calibration is/are calibrated by using reference standards traceable to National/International standards. | | | | | | | | | | | |
| The calibration results reported are valid at the time of and under the conditions of measurement. | | | | | | | | | | | |
| This report may not be reproduced, except in full, without the prior permission of "Pin Point Engineers". | | | | | | | | | | | |
| Any correction in this certificate invalidated the certificate. | | | | | | | | | | | |
|  |  |  |  |  | ***CALIBRATION CERTIFICATE NO: PPE078SW007201/003*** | | | | | | |
|  |  |  |  |  |  |  |  |  |  |  |  |
| **CLIENT NAME :** | | | **Rusan Pharma Limited** | | | | | | | |  |
| **ADDRESS:** | |  | Khasra No.: 122 MI, Central Hope Town, Behind Pharma City, Selaqui, Dehradun-248197, INDIA | | | | | | | | |
|  |  |  |
| **ENVIRONMENTAL CONDITION:** | | | **AMBIENT TEMPERATURE & HUMIDITY** | | | | |  |  |  |  |
| ***INSTRUMENT DETAILS*** | | | | | | | | | | | |
| **INSTRUMENT** | | Stop Watch | | | | **DEPARTMENT** | | Quality Control | | | |
| **MAKE** | | Racer | | | | **CAL. RANGFE** | | 2 Hours | | | |
| **RANGE** | | 0 – 23 Min : 59 Sec | | | | **LOCATION** | | Q.C. Lab | | | |
| **ACCURACY** | | ± 1% of Full Scale | | | | **TAG ID** | | SW/QCE/003 | | | |
| **LEAST COUNT** | | 1 Sec. | | | | **CAL. DATE** | | 09/07/2021 | | | |
| **UNIT OF MEASURMENT** | | Sec. | | |  | **DUE ON** | | 08/07/2022 | | | |
| **CALIBRATION PROCEDURE NUMBER** | | | | | | PPE/SOP/TIME/TM/01-00 | | | | | |
| **OBSERVATION TABLE** | | | | | | | | | | | |
| **Sr. No** | **STANDARD READINGS ON UNIT UNDER CALIBRATION (HH:MM:SS: 1/100s)** | | | **OBSERVED READINGS ON STANDARD INSTRUMENT (HH:MM:SS:1/100s)** | | | | **ERROR (HH:MM:SS:1/100s)** | | | |
|
|
| **1** | 0:00:00:00 | | | 0:00:00:01 | | | | 0:00:00:01 | | | |
| **2** | 0:30:00:00 | | | 0:30:00:02 | | | | 0:00:00:02 | | | |
| **3** | 1:00:00 | | | 1:00:00 | | | | 0:00:00 | | | |
| **4** | 2:00:00 | | | 2:00:00 | | | | 0:00:00 | | | |
| **Note: (1/100 s  reading is available till 1 Hours)** | | | | | | | | | | | |
| **REMARK: THE STOP WATCH IS CALIBRATED SUCCESSFULLY WITHIN THE SPECIFIED LIMITS** | | | | | | | | | | | |
|  |  |  |  |  |  |  |  |  |  |  |  |
| ***REFERENCE STANDARDS USED*** | | | | | | | | | | | |
| **INSTRUMENT** | | Digital Stop Watch | | | | **RANGE** | | 0-24 Hour | | | |
| **MAKE** | | Racer | | | | **REPORT NO** | | 2021/01/0061 | |  |  |
| **INST ID. NO** | | PPE/C/SW/01 | | | |  |  | **(Traceable To NABL)** | |  | |
|  |  |  |  |  |  | **DUE DATE** | | 06/01/2022 | |  |  |
|  | | | |  | | | |  | | | |
|
|
|
| **CALIBRATED BY** | | | | **CHECKED & APPROVED BY** | | | | **REVIEWED BY** | | | |
| ***NOTE*** | |  |  |  |  |  |  |  |  |  |  |
| This report refers only to the particular item calibrated at site/laboratary. | | | | | | | | | | |  |
| Pin Point Engineers is not liable for any change in calibration data & performance specifications on account of malfunctioning the standards/instrument/equipment covered by this report due to damage caused to it after issuance of this report. | | | | | | | | | | | |
|
| The Standard(s) used for calibration is/are calibrated by using reference standards traceable to National/International standards. | | | | | | | | | | | |
| The calibration results reported are valid at the time of and under the conditions of measurement. | | | | | | | | | | | |
| This report may not be reproduced, except in full, without the prior permission of "Pin Point Engineers". | | | | | | | | | | | |
| Any correction in this certificate invalidated the certificate. | | | | | | | | | | | |