## Hygrometer
CALIBRATION CERTIFICATE NO: PPE078DTHM1121/001
Client Name: |  |  | Rusan Pharma Limited
ADDRESS: |  |  | Khasra No.: 122 MI, Central Hope Town, Behind Pharma City, Selaqui,
Dehradun-248197, INDIA
ENVIRONMENTAL CONDITION: |  |  | AMBIENT TEMPERATURE & HUMIDITY
INSTRUMENT DETAILS
INSTRUMENT |  |  | Digital Hygrometer |  |  | DEPARTMENT |  | QC
MAKE |  |  | Rotronic
RANGE |  |  | -10 to 50°C/ 10 to 99% RH |  |  | LOCATION |  | Chemical Testing Lab
ACCURACY |  |  | ± 1°C / ± 5% RH |  |  | TAG NO |  | DTHM-121 |  | print
LEAST COUNT |  |  | 0.1°C & 0.1% RH |  |  | CAL. DATE |  | 2021-11-02
UNIT OF MEASUREMENT |  |  | %RH & °C |  |  | DUE ON |  | 2022-11-01
CALIBRATION PROCEDURE NUMBER | [5 empty] | PPE/SOP/TEMP/DTH/01-00
OBSERVATION TABLE
Sr. No | STANDARD READINGS ON UNIT UNDER CALIBRATION |  |  |  | OBSERVED READING ON MASTER INSTRUMENT |  |  | ERROR
| RH |  | TEMP. |  | RH | TEMP. |  | RH | TEMP.
1 | 40 |  | 10 |  | 40.3 | 10.1 |  | 0.3 | 0.1
2 | 60 |  | 30 |  | 60.3 | 30.2 |  | 0.3 | 0.2
3 | 80 |  | 40 |  | 80.2 | 40.3 |  | 0.2 | 0.3
REMARK: THE DIGITAL THERMO HYGROMETER IS CALIBRATED SUCCESSFULLY WITHIN THE SPECIFIED LIMITS
REFERENCE STANDARDS USED
INSTRUMENT |  | Temperature & Humidity Controller |  |  |  | REPORT NO |  | 2021/01/0059
MAKE |  | Ambetronics/ (TRH-307) |  |  |  | DUE DATE |  | 2022-01-06
INST ID. NO |  | PPE/C/TEMP/THC/01 |  |  |  | ACCURACY |  | ± 0.3 % F.S / ±1.8% RH
RANGE |  | (-40) to 120 °C / 0-100%RH |  |  |  | LEAST COUNT |  | 0.1°C / 0.1% RH
CALIBRATED BY |  |  |  | CHECKED & APPROVED BY |  |  | REVIEWED BY
NOTE
This report refers only to the particular item calibrated at site/laboratary.
Pin Point Engineers is not liable for any change in calibration data & performance specifications on account of malfunctioning the standards/instrument/equipment covered by this report due to damage caused to it after issuance of this report.
The Standard(s) used for calibration is/are calibrated by using reference standards traceable to National/International standards.
The calibration results reported are valid at the time of and under the conditions of measurement.
This report may not be reproduced, except in full, without the prior permission of "Pin Point Engineers".
Any correction in this certificate invalidated the certificate.
CALIBRATION CERTIFICATE NO: PPE078DTHM1121/002
Client Name: |  |  | Rusan Pharma Limited
ADDRESS: |  |  | Khasra No.: 122 MI, Central Hope Town, Behind Pharma City, Selaqui,
Dehradun-248197, INDIA
ENVIRONMENTAL CONDITION: |  |  | AMBIENT TEMPERATURE & HUMIDITY
INSTRUMENT DETAILS
INSTRUMENT |  |  | Digital Hygrometer |  |  | DEPARTMENT |  | QC
MAKE |  |  | Rotronic
RANGE |  |  | -10 to 50°C/ 10 to 99% RH |  |  | LOCATION |  | Instrument Lab
ACCURACY |  |  | ± 1°C / ± 5% RH |  |  | TAG NO |  | DTHM-116
LEAST COUNT |  |  | 0.1°C & 0.1% RH |  |  | CAL. DATE |  | 2021-11-02
UNIT OF MEASUREMENT |  |  | %RH & °C |  |  | DUE ON |  | 2022-11-01
CALIBRATION PROCEDURE NUMBER | [5 empty] | PPE/SOP/TEMP/DTH/01-00
OBSERVATION TABLE
Sr. No | STANDARD READINGS ON UNIT UNDER CALIBRATION |  |  |  | OBSERVED READING ON MASTER INSTRUMENT |  |  | ERROR
| RH |  | TEMP. |  | RH | TEMP. |  | RH | TEMP.
1 | 40 |  | 10 |  | 40.2 | 10.2 |  | 0.2 | 0.2
2 | 60 |  | 30 |  | 60.4 | 30.3 |  | 0.4 | 0.3
3 | 80 |  | 40 |  | 80.3 | 40.4 |  | 0.3 | 0.4
REMARK: THE DIGITAL THERMO HYGROMETER IS CALIBRATED SUCCESSFULLY WITHIN THE SPECIFIED LIMITS
REFERENCE STANDARDS USED
INSTRUMENT |  | Temperature & Humidity Controller |  |  |  | REPORT NO |  | 2021/01/0059
MAKE |  | Ambetronics/ (TRH-307) |  |  |  | DUE DATE |  | 2022-01-06
INST ID. NO |  | PPE/C/TEMP/THC/01 |  |  |  | ACCURACY |  | ± 0.3 % F.S / ±1.8% RH
RANGE |  | (-40) to 120 °C / 0-100%RH |  |  |  | LEAST COUNT |  | 0.1°C / 0.1% RH
CALIBRATED BY |  |  |  | CHECKED & APPROVED BY |  |  | REVIEWED BY
NOTE
This report refers only to the particular item calibrated at site/laboratary.
Pin Point Engineers is not liable for any change in calibration data & performance specifications on account of malfunctioning the standards/instrument/equipment covered by this report due to damage caused to it after issuance of this report.
The Standard(s) used for calibration is/are calibrated by using reference standards traceable to National/International standards.
The calibration results reported are valid at the time of and under the conditions of measurement.
This report may not be reproduced, except in full, without the prior permission of "Pin Point Engineers".
Any correction in this certificate invalidated the certificate.
CALIBRATION CERTIFICATE NO: PPE078DTHM1121/003
Client Name: |  |  | Rusan Pharma Limited
ADDRESS: |  |  | Khasra No.: 122 MI, Central Hope Town, Behind Pharma City, Selaqui,
Dehradun-248197, INDIA
ENVIRONMENTAL CONDITION: |  |  | AMBIENT TEMPERATURE & HUMIDITY
INSTRUMENT DETAILS
INSTRUMENT |  |  | Digital Hygrometer |  |  | DEPARTMENT |  | QC
MAKE |  |  | Rotronic
RANGE |  |  | -10 to 50°C/ 10 to 99% RH |  |  | LOCATION |  | Instrument Lab
ACCURACY |  |  | ± 1°C / ± 5% RH |  |  | TAG NO |  | DTHM-110
LEAST COUNT |  |  | 0.1°C & 0.1% RH |  |  | CAL. DATE |  | 2021-11-02
UNIT OF MEASUREMENT |  |  | %RH & °C |  |  | DUE ON |  | 2022-11-01
CALIBRATION PROCEDURE NUMBER | [5 empty] | PPE/SOP/TEMP/DTH/01-00
OBSERVATION TABLE
Sr. No | STANDARD READINGS ON UNIT UNDER CALIBRATION |  |  |  | OBSERVED READING ON MASTER INSTRUMENT |  |  | ERROR
| RH |  | TEMP. |  | RH | TEMP. |  | RH | TEMP.
1 | 40 |  | 10 |  | 40.2 | 10.1 |  | 0.2 | 0.1
2 | 60 |  | 30 |  | 60.3 | 30.2 |  | 0.3 | 0.2
3 | 80 |  | 40 |  | 80.3 | 40.3 |  | 0.3 | 0.3
REMARK: THE DIGITAL THERMO HYGROMETER IS CALIBRATED SUCCESSFULLY WITHIN THE SPECIFIED LIMITS
REFERENCE STANDARDS USED
INSTRUMENT |  | Temperature & Humidity Controller |  |  |  | REPORT NO |  | 2021/01/0059
MAKE |  | Ambetronics/ (TRH-307) |  |  |  | DUE DATE |  | 2022-01-06
INST ID. NO |  | PPE/C/TEMP/THC/01 |  |  |  | ACCURACY |  | ± 0.3 % F.S / ±1.8% RH
RANGE |  | (-40) to 120 °C / 0-100%RH |  |  |  | LEAST COUNT |  | 0.1°C / 0.1% RH
CALIBRATED BY |  |  |  | CHECKED & APPROVED BY |  |  | REVIEWED BY
NOTE
This report refers only to the particular item calibrated at site/laboratary.
Pin Point Engineers is not liable for any change in calibration data & performance specifications on account of malfunctioning the standards/instrument/equipment covered by this report due to damage caused to it after issuance of this report.
The Standard(s) used for calibration is/are calibrated by using reference standards traceable to National/International standards.
The calibration results reported are valid at the time of and under the conditions of measurement.
This report may not be reproduced, except in full, without the prior permission of "Pin Point Engineers".
Any correction in this certificate invalidated the certificate.
CALIBRATION CERTIFICATE NO: PPE078DTHM1121/004
Client Name: |  |  | Rusan Pharma Limited
ADDRESS: |  |  | Khasra No.: 122 MI, Central Hope Town, Behind Pharma City, Selaqui,
Dehradun-248197, INDIA
ENVIRONMENTAL CONDITION: |  |  | AMBIENT TEMPERATURE & HUMIDITY
INSTRUMENT DETAILS
INSTRUMENT |  |  | Digital Hygrometer |  |  | DEPARTMENT |  | QC
MAKE |  |  | HTC
RANGE |  |  | -10 to 50°C/ 10 to 99% RH |  |  | LOCATION |  | Instrument Lab
ACCURACY |  |  | ± 1°C / ± 5% RH |  |  | TAG NO |  | DTHM-041
LEAST COUNT |  |  | 0.1°C & 1% RH |  |  | CAL. DATE |  | 2021-11-02
UNIT OF MEASUREMENT |  |  | %RH & °C |  |  | DUE ON |  | 2022-11-01
CALIBRATION PROCEDURE NUMBER | [5 empty] | PPE/SOP/TEMP/DTH/01-00
OBSERVATION TABLE
Sr. No | STANDARD READINGS ON UNIT UNDER CALIBRATION |  |  |  | OBSERVED READING ON MASTER INSTRUMENT |  |  | ERROR
| RH |  | TEMP. |  | RH | TEMP. |  | RH | TEMP.
1 | 40 |  | 10 |  | 40.2 | 10.2 |  | 0.2 | 0.2
2 | 60 |  | 30 |  | 60.3 | 30.2 |  | 0.3 | 0.2
3 | 80 |  | 40 |  | 80.3 | 40.3 |  | 0.3 | 0.3
REMARK: THE DIGITAL THERMO HYGROMETER IS CALIBRATED SUCCESSFULLY WITHIN THE SPECIFIED LIMITS
REFERENCE STANDARDS USED
INSTRUMENT |  | Temperature & Humidity Controller |  |  |  | REPORT NO |  | 2021/01/0059
MAKE |  | Ambetronics/ (TRH-307) |  |  |  | DUE DATE |  | 2022-01-06
INST ID. NO |  | PPE/C/TEMP/THC/01 |  |  |  | ACCURACY |  | ± 0.3 % F.S / ±1.8% RH
RANGE |  | (-40) to 120 °C / 0-100%RH |  |  |  | LEAST COUNT |  | 0.1°C / 0.1% RH
CALIBRATED BY |  |  |  | CHECKED & APPROVED BY |  |  | REVIEWED BY
NOTE
This report refers only to the particular item calibrated at site/laboratary.
Pin Point Engineers is not liable for any change in calibration data & performance specifications on account of malfunctioning the standards/instrument/equipment covered by this report due to damage caused to it after issuance of this report.
The Standard(s) used for calibration is/are calibrated by using reference standards traceable to National/International standards.
The calibration results reported are valid at the time of and under the conditions of measurement.
This report may not be reproduced, except in full, without the prior permission of "Pin Point Engineers".
Any correction in this certificate invalidated the certificate.
CALIBRATION CERTIFICATE NO: PPE078DTHM1121/005
Client Name: |  |  | Rusan Pharma Limited
ADDRESS: |  |  | Khasra No.: 122 MI, Central Hope Town, Behind Pharma City, Selaqui,
Dehradun-248197, INDIA
ENVIRONMENTAL CONDITION: |  |  | AMBIENT TEMPERATURE & HUMIDITY
INSTRUMENT DETAILS
INSTRUMENT |  |  | Digital Hygrometer |  |  | DEPARTMENT |  | QC
MAKE |  |  | HTC
RANGE |  |  | -10 to 50°C/ 10 to 99% RH |  |  | LOCATION |  | Instrument Lab
ACCURACY |  |  | ± 1°C / ± 5% RH |  |  | TAG NO |  | DTHM-101
LEAST COUNT |  |  | 0.1°C & 1% RH |  |  | CAL. DATE |  | 2021-11-02
UNIT OF MEASUREMENT |  |  | %RH & °C |  |  | DUE ON |  | 2022-11-01
CALIBRATION PROCEDURE NUMBER | [5 empty] | PPE/SOP/TEMP/DTH/01-00
OBSERVATION TABLE
Sr. No | STANDARD READINGS ON UNIT UNDER CALIBRATION |  |  |  | OBSERVED READING ON MASTER INSTRUMENT |  |  | ERROR
| RH |  | TEMP. |  | RH | TEMP. |  | RH | TEMP.
1 | 40 |  | 10 |  | 40.1 | 10.2 |  | 0.1 | 0.2
2 | 60 |  | 30 |  | 60.2 | 30.2 |  | 0.2 | 0.2
3 | 80 |  | 40 |  | 80.2 | 40.3 |  | 0.2 | 0.3
REMARK: THE DIGITAL THERMO HYGROMETER IS CALIBRATED SUCCESSFULLY WITHIN THE SPECIFIED LIMITS
REFERENCE STANDARDS USED
INSTRUMENT |  | Temperature & Humidity Controller |  |  |  | REPORT NO |  | 2021/01/0059
MAKE |  | Ambetronics/ (TRH-307) |  |  |  | DUE DATE |  | 2022-01-06
INST ID. NO |  | PPE/C/TEMP/THC/01 |  |  |  | ACCURACY |  | ± 0.3 % F.S / ±1.8% RH
RANGE |  | (-40) to 120 °C / 0-100%RH |  |  |  | LEAST COUNT |  | 0.1°C / 0.1% RH
CALIBRATED BY |  |  |  | CHECKED & APPROVED BY |  |  | REVIEWED BY
NOTE
This report refers only to the particular item calibrated at site/laboratary.
Pin Point Engineers is not liable for any change in calibration data & performance specifications on account of malfunctioning the standards/instrument/equipment covered by this report due to damage caused to it after issuance of this report.
The Standard(s) used for calibration is/are calibrated by using reference standards traceable to National/International standards.
The calibration results reported are valid at the time of and under the conditions of measurement.
This report may not be reproduced, except in full, without the prior permission of "Pin Point Engineers".
Any correction in this certificate invalidated the certificate.
CALIBRATION CERTIFICATE NO: PPE078DTHM1121/006
Client Name: |  |  | Rusan Pharma Limited
ADDRESS: |  |  | Khasra No.: 122 MI, Central Hope Town, Behind Pharma City, Selaqui,
Dehradun-248197, INDIA
ENVIRONMENTAL CONDITION: |  |  | AMBIENT TEMPERATURE & HUMIDITY
INSTRUMENT DETAILS
INSTRUMENT |  |  | Digital Hygrometer |  |  | DEPARTMENT |  | QC
MAKE |  |  | HTC
RANGE |  |  | -10 to 50°C/ 10 to 99% RH |  |  | LOCATION |  | Instrument Lab
ACCURACY |  |  | ± 1°C / ± 5% RH |  |  | TAG NO |  | DTHM-003
LEAST COUNT |  |  | 0.1°C & 1% RH |  |  | CAL. DATE |  | 2021-11-02
UNIT OF MEASUREMENT |  |  | %RH & °C |  |  | DUE ON |  | 2022-11-01
CALIBRATION PROCEDURE NUMBER | [5 empty] | PPE/SOP/TEMP/DTH/01-00
OBSERVATION TABLE
Sr. No | STANDARD READINGS ON UNIT UNDER CALIBRATION |  |  |  | OBSERVED READING ON MASTER INSTRUMENT |  |  | ERROR
| RH |  | TEMP. |  | RH | TEMP. |  | RH | TEMP.
1 | 40 |  | 10 |  | 40.2 | 10.1 |  | 0.2 | 0.1
2 | 60 |  | 30 |  | 60.3 | 30.2 |  | 0.3 | 0.2
3 | 80 |  | 40 |  | 80.3 | 40.3 |  | 0.3 | 0.3
REMARK: THE DIGITAL THERMO HYGROMETER IS CALIBRATED SUCCESSFULLY WITHIN THE SPECIFIED LIMITS
REFERENCE STANDARDS USED
INSTRUMENT |  | Temperature & Humidity Controller |  |  |  | REPORT NO |  | 2021/01/0059
MAKE |  | Ambetronics/ (TRH-307) |  |  |  | DUE DATE |  | 2022-01-06
INST ID. NO |  | PPE/C/TEMP/THC/01 |  |  |  | ACCURACY |  | ± 0.3 % F.S / ±1.8% RH
RANGE |  | (-40) to 120 °C / 0-100%RH |  |  |  | LEAST COUNT |  | 0.1°C / 0.1% RH
CALIBRATED BY |  |  |  | CHECKED & APPROVED BY |  |  | REVIEWED BY
NOTE
This report refers only to the particular item calibrated at site/laboratary.
Pin Point Engineers is not liable for any change in calibration data & performance specifications on account of malfunctioning the standards/instrument/equipment covered by this report due to damage caused to it after issuance of this report.
The Standard(s) used for calibration is/are calibrated by using reference standards traceable to National/International standards.
The calibration results reported are valid at the time of and under the conditions of measurement.
This report may not be reproduced, except in full, without the prior permission of "Pin Point Engineers".
Any correction in this certificate invalidated the certificate.
CALIBRATION CERTIFICATE NO: PPE078DTHM1121/030
CLIENT NAME : |  |  | Rusan Pharma Limited
ADDRESS: |  |  | Khasra No.: 122 MI, Central Hope Town, Behind Pharma City, Selaqui,
Dehradun-248197, INDIA
ENVIRONMENTAL CONDITION: |  |  | AMBIENT TEMPERATURE & HUMIDITY
INSTRUMENT DETAILS
INSTRUMENT |  |  | Digital Thermo Hygrometer |  |  | DEPARTMENT |  | QC
MAKE |  |  | HTC
RANGE |  |  | -10 to 50°C/ 10 to 99% RH |  |  | LOCATION |  | Micro
ACCURACY |  |  | ± 1°C / ± 5% RH |  |  | TAG NO |  | DTHM-042
LEAST COUNT |  |  | 0.1°C & 1% RH |  |  | CAL. DATE |  | 2021-11-11
UNIT OF MEASUREMENT |  |  | %RH & °C |  |  | DUE ON |  | 2022-11-10
CALIBRATION PROCEDURE NUMBER | [5 empty] | PPE/SOP/TEMP/DTH/01-00
OBSERVATION TABLE
Sr. No | STANDARD READINGS ON UNIT UNDER CALIBRATION |  |  |  | OBSERVED READING ON MASTER INSTRUMENT |  |  | ERROR (±)
| RH |  | TEMP. |  | RH | TEMP. |  | RH | TEMP.
1 | 40 |  | 10 |  | 40.2 | 10.1 |  | 0.2 | 0.1
2 | 60 |  | 30 |  | 60.3 | 30.1 |  | 0.3 | 0.1
3 | 80 |  | 40 |  | 80.4 | 40.1 |  | 0.4 | 0.1
REMARK: THE DIGITAL THERMO HYGROMETER IS CALIBRATED SUCCESSFULLY WITHIN THE SPECIFIED LIMITS
REFERENCE STANDARDS USED
INSTRUMENT |  | Temperature & Humidity Controller |  |  |  | REPORT NO |  | 2021/01/0059
MAKE |  | Ambetronics/ (TRH-307) | [5 empty] | (Traceable To NABL)
INST ID. NO |  | PPE/C/TEMP/THC/01 |  |  |  | DUE DATE |  | 2022-01-06
RANGE |  | (-40) to 120 °C / 0-100%RH |  |  |  | ACCURACY |  | ± 0.3 % F.S / ±1.8% RH
[6 empty] | LEAST COUNT |  | 0.1°C / 0.1% RH
CALIBRATED BY |  |  |  | CHECKED & APPROVED BY |  |  | REVIEWED BY
NOTE
This report refers only to the particular item calibrated at site/laboratary.
Pin Point Engineers is not liable for any change in calibration data & performance specifications on account of malfunctioning the standards/instrument/equipment covered by this report due to damage caused to it after issuance of this report.
The Standard(s) used for calibration is/are calibrated by using reference standards traceable to National/International standards.
The calibration results reported are valid at the time of and under the conditions of measurement.
This report may not be reproduced, except in full, without the prior permission of "Pin Point Engineers".
Any correction in this certificate invalidated the certificate.
CALIBRATION CERTIFICATE NO: PPE078DTHM1121/031
CLIENT NAME: |  |  | Rusan Pharma Limited
ADDRESS: |  |  | Khasra No.: 122 MI, Central Hope Town, Behind Pharma City, Selaqui,
Dehradun-248197, INDIA
ENVIRONMENTAL CONDITION: |  |  | AMBIENT TEMPERATURE & HUMIDITY
INSTRUMENT DETAILS
INSTRUMENT |  |  | Digital Thermo Hygrometer |  |  | DEPARTMENT |  | QC
MAKE |  |  | Rotronic
RANGE |  |  | -10 to 50°C/ 10 to 99% RH |  |  | LOCATION |  | Micro
ACCURACY |  |  | ± 1°C / ± 5% RH |  |  | TAG NO |  | DTHM-113
LEAST COUNT |  |  | 0.1°C & 1% RH |  |  | CAL. DATE |  | 2021-11-11
UNIT OF MEASUREMENT |  |  | %RH & °C |  |  | DUE ON |  | 2022-11-10
CALIBRATION PROCEDURE NUMBER | [5 empty] | PPE/SOP/TEMP/DTH/01-00
OBSERVATION TABLE
Sr. No | STANDARD READINGS ON UNIT UNDER CALIBRATION |  |  |  | OBSERVED READING ON MASTER INSTRUMENT |  |  | ERROR
| RH |  | TEMP. |  | RH | TEMP. |  | RH | TEMP.
1 | 40 |  | 10 |  | 40.2 | 10.1 |  | 0.2 | 0.1
2 | 60 |  | 30 |  | 60.3 | 30.2 |  | 0.3 | 0.2
3 | 80 |  | 40 |  | 80.4 | 40.2 |  | 0.4 | 0.2
REMARK: THE DIGITAL THERMO HYGROMETER IS CALIBRATED SUCCESSFULLY WITHIN THE SPECIFIED LIMITS
REFERENCE STANDARDS USED
INSTRUMENT |  | Temperature & Humidity Controller |  |  |  | REPORT NO |  | 2021/01/0059
MAKE |  | Ambetronics/ (TRH-307) |  |  |  | DUE DATE |  | 2022-01-06
INST ID. NO |  | PPE/C/TEMP/THC/01 |  |  |  | ACCURACY |  | ± 0.3 % F.S / ±1.8% RH
RANGE |  | (-40) to 120 °C / 0-100%RH |  |  |  | LEAST COUNT |  | 0.1°C / 0.1% RH
CALIBRATED BY |  |  |  | CHECKED & APPROVED BY |  |  | REVIEWED BY
NOTE
This report refers only to the particular item calibrated at site/laboratary.
Pin Point Engineers is not liable for any change in calibration data & performance specifications on account of malfunctioning the standards/instrument/equipment covered by this report due to damage caused to it after issuance of this report.
The Standard(s) used for calibration is/are calibrated by using reference standards traceable to National/International standards.
The calibration results reported are valid at the time of and under the conditions of measurement.
This report may not be reproduced, except in full, without the prior permission of "Pin Point Engineers".
Any correction in this certificate invalidated the certificate.
CALIBRATION CERTIFICATE NO: PPE078DTHM1121/032
CLIENT NAME: |  |  | Rusan Pharma Limited
ADDRESS: |  |  | Khasra No.: 122 MI, Central Hope Town, Behind Pharma City, Selaqui,
Dehradun-248197, INDIA
ENVIRONMENTAL CONDITION: |  |  | AMBIENT TEMPERATURE & HUMIDITY
INSTRUMENT DETAILS
INSTRUMENT |  |  | Digital Thermo Hygrometer |  |  | DEPARTMENT |  | QC
MAKE |  |  | Instrumentics
RANGE |  |  | -10 to 50°C/ 10 to 99% RH |  |  | LOCATION |  | Micro
ACCURACY |  |  | ± 1°C / ± 5% RH |  |  | TAG NO |  | DTHM-001
LEAST COUNT |  |  | 0.1°C & 1% RH |  |  | CAL. DATE |  | 2021-11-11
UNIT OF MEASUREMENT |  |  | %RH & °C |  |  | DUE ON |  | 2022-11-10
CALIBRATION PROCEDURE NUMBER | [5 empty] | PPE/SOP/TEMP/DTH/01-00
OBSERVATION TABLE
Sr. No | STANDARD READINGS ON UNIT UNDER CALIBRATION |  |  |  | OBSERVED READING ON MASTER INSTRUMENT |  |  | ERROR(±)
| RH |  | TEMP. |  | RH | TEMP. |  | RH | TEMP.
1 | 40 |  | 10 |  | 40.3 | 10.2 |  | 0.3 | 0.2
2 | 60 |  | 30 |  | 60.4 | 30.2 |  | 0.4 | 0.2
3 | 80 |  | 40 |  | 80.4 | 40.2 |  | 0.4 | 0.2
REMARK: THE DIGITAL THERMO HYGROMETER IS CALIBRATED SUCCESSFULLY WITHIN THE SPECIFIED LIMITS
REFERENCE STANDARDS USED
INSTRUMENT |  | Temperature & Humidity Controller |  |  |  | REPORT NO |  | 2021/01/0059
MAKE |  | Ambetronics/ (TRH-307) | [5 empty] | (Traceable to NABL)
INST ID. NO |  | PPE/C/TEMP/THC/01 |  |  |  | DUE DATE |  | 2022-01-06
RANGE |  | (-40) to 120 °C / 0-100%RH |  |  |  | ACCURACY |  | ± 0.3 % F.S / ±1.8% RH
[6 empty] | LEAST COUNT |  | 0.1°C / 0.1% RH
CALIBRATED BY |  |  |  | CHECKED & APPROVED BY |  |  | REVIEWED BY
NOTE
This report refers only to the particular item calibrated at site/laboratary.
Pin Point Engineers is not liable for any change in calibration data & performance specifications on account of malfunctioning the standards/instrument/equipment covered by this report due to damage caused to it after issuance of this report.
The Standard(s) used for calibration is/are calibrated by using reference standards traceable to National/International standards.
The calibration results reported are valid at the time of and under the conditions of measurement.
This report may not be reproduced, except in full, without the prior permission of "Pin Point Engineers".
Any correction in this certificate invalidated the certificate.
CALIBRATION CERTIFICATE NO: PPE078DTHM1121/033
CLIENT NAME: |  |  | Rusan Pharma Limited
ADDRESS: |  |  | Khasra No.: 122 MI, Central Hope Town, Behind Pharma City, Selaqui,
Dehradun-248197, INDIA
ENVIRONMENTAL CONDITION: |  |  | AMBIENT TEMPERATURE & HUMIDITY
INSTRUMENT DETAILS
INSTRUMENT |  |  | Digital Thermo Hygrometer |  |  | DEPARTMENT |  | QC
MAKE |  |  | HTC
RANGE |  |  | -10 to 50°C/ 10 to 99% RH |  |  | LOCATION |  | MLT Room
ACCURACY |  |  | ± 1°C / ± 5% RH |  |  | TAG NO |  | DTHM-002
LEAST COUNT |  |  | 0.1°C & 1% RH |  |  | CAL. DATE |  | 2021-11-02
UNIT OF MEASUREMENT |  |  | %RH & °C |  |  | DUE ON |  | 2022-11-01
CALIBRATION PROCEDURE NUMBER | [5 empty] | PPE/SOP/TEMP/DTH/01-00
OBSERVATION TABLE
Sr. No | STANDARD READINGS ON UNIT UNDER CALIBRATION |  |  |  | OBSERVED READING ON MASTER INSTRUMENT |  |  | ERROR (±)
| RH |  | TEMP. |  | RH | TEMP. |  | RH | TEMP.
1 | 40 |  | 10 |  | 40.1 | 10.1 |  | 0.1 | 0.1
2 | 60 |  | 30 |  | 60.2 | 30.1 |  | 0.2 | 0.1
3 | 80 |  | 40 |  | 80.4 | 40.2 |  | 0.4 | 0.2
REMARK: THE DIGITAL THERMO HYGROMETER IS CALIBRATED SUCCESSFULLY WITHIN THE SPECIFIED LIMITS
REFERENCE STANDARDS USED
INSTRUMENT |  | Temperature & Humidity Controller |  |  |  | REPORT NO |  | 2021/01/0059
(Traceable To NABL)
MAKE |  | Ambetronics/ (TRH-307) |  |  |  | DUE DATE |  | 2022-01-06
INST ID. NO |  | PPE/C/TEMP/THC/01 |  |  |  | ACCURACY |  | ± 0.3 % F.S / ±1.8% RH
RANGE |  | (-40) to 120 °C / 0-100%RH |  |  |  | LEAST COUNT |  | 0.1°C / 0.1% RH
CALIBRATED BY |  |  |  | CHECKED & APPROVED BY |  |  | REVIEWED BY
NOTE
This report refers only to the particular item calibrated at site/laboratary.
Pin Point Engineers is not liable for any change in calibration data & performance specifications on account of malfunctioning the standards/instrument/equipment covered by this report due to damage caused to it after issuance of this report.
The Standard(s) used for calibration is/are calibrated by using reference standards traceable to National/International standards.
The calibration results reported are valid at the time of and under the conditions of measurement.
This report may not be reproduced, except in full, without the prior permission of "Pin Point Engineers".
Any correction in this certificate invalidated the certificate.
CALIBRATION CERTIFICATE NO: PPE078DTHM1121/034
CLIENT NAME: |  |  | Rusan Pharma Limited
ADDRESS: |  |  | Khasra No.: 122 MI, Central Hope Town, Behind Pharma City, Selaqui,
Dehradun-248197, INDIA
ENVIRONMENTAL CONDITION: |  |  | AMBIENT TEMPERATURE & HUMIDITY
INSTRUMENT DETAILS
INSTRUMENT |  |  | Digital Thermo Hygrometer |  |  | DEPARTMENT |  | QC
MAKE |  |  | HTC
RANGE |  |  | -10 to 50°C/ 10 to 99% RH |  |  | LOCATION |  | Instrument Lab-II
ACCURACY |  |  | ± 1°C / ± 5% RH |  |  | TAG NO |  | DTHM-004
LEAST COUNT |  |  | 0.1°C & 1% RH |  |  | CAL. DATE |  | 2021-11-02
UNIT OF MEASUREMENT |  |  | %RH & °C |  |  | DUE ON |  | 2022-11-01
CALIBRATION PROCEDURE NUMBER | [5 empty] | PPE/SOP/TEMP/DTH/01-00
OBSERVATION TABLE
Sr. No | STANDARD READINGS ON UNIT UNDER CALIBRATION |  |  |  | OBSERVED READING ON MASTER INSTRUMENT |  |  | ERROR (±)
| RH |  | TEMP. |  | RH | TEMP. |  | RH | TEMP.
1 | 40 |  | 10 |  | 40.1 | 10.1 |  | 0.1 | 0.1
2 | 60 |  | 30 |  | 60.1 | 30.1 |  | 0.1 | 0.1
3 | 80 |  | 40 |  | 80.3 | 40.2 |  | 0.3 | 0.2
REMARK: THE DIGITAL THERMO HYGROMETER IS CALIBRATED SUCCESSFULLY WITHIN THE SPECIFIED LIMITS
REFERENCE STANDARDS USED
INSTRUMENT |  | Temperature & Humidity Controller |  |  |  | REPORT NO |  | 2021/01/0059
(Traceable To NABL)
MAKE |  | Ambetronics/ (TRH-307) |  |  |  | DUE DATE |  | 2022-01-06
INST ID. NO |  | PPE/C/TEMP/THC/01 |  |  |  | ACCURACY |  | ± 0.3 % F.S / ±1.8% RH
RANGE |  | (-40) to 120 °C / 0-100%RH |  |  |  | LEAST COUNT |  | 0.1°C / 0.1% RH
CALIBRATED BY |  |  |  | CHECKED & APPROVED BY |  |  | REVIEWED BY
NOTE
This report refers only to the particular item calibrated at site/laboratary.
Pin Point Engineers is not liable for any change in calibration data & performance specifications on account of malfunctioning the standards/instrument/equipment covered by this report due to damage caused to it after issuance of this report.
The Standard(s) used for calibration is/are calibrated by using reference standards traceable to National/International standards.
The calibration results reported are valid at the time of and under the conditions of measurement.
This report may not be reproduced, except in full, without the prior permission of "Pin Point Engineers".
Any correction in this certificate invalidated the certificate.
CALIBRATION CERTIFICATE NO: PPE078DTHM1121/035
CLIENT NAME: |  |  | Rusan Pharma Limited
ADDRESS: |  |  | Khasra No.: 122 MI, Central Hope Town, Behind Pharma City, Selaqui,
Dehradun-248197, INDIA
ENVIRONMENTAL CONDITION: |  |  | AMBIENT TEMPERATURE & HUMIDITY
INSTRUMENT DETAILS
INSTRUMENT |  |  | Digital Thermo Hygrometer |  |  | DEPARTMENT |  | QC
MAKE |  |  | HTC
RANGE |  |  | -10 to 50°C/ 10 to 99% RH |  |  | LOCATION |  | Cooling Zone
ACCURACY |  |  | ± 1°C / ± 5% RH |  |  | TAG NO |  | DTHM-005
LEAST COUNT |  |  | 0.1°C & 1% RH |  |  | CAL. DATE |  | 2021-11-02
UNIT OF MEASUREMENT |  |  | %RH & °C |  |  | DUE ON |  | 2022-11-01
CALIBRATION PROCEDURE NUMBER | [5 empty] | PPE/SOP/TEMP/DTH/01-00
OBSERVATION TABLE
Sr. No | STANDARD READINGS ON UNIT UNDER CALIBRATION |  |  |  | OBSERVED READING ON MASTER INSTRUMENT |  |  | ERROR (±)
| RH |  | TEMP. |  | RH | TEMP. |  | RH | TEMP.
1 | 40 |  | 10 |  | 40.1 | 10.1 |  | 0.1 | 0.1
2 | 60 |  | 30 |  | 60.2 | 30.1 |  | 0.2 | 0.1
3 | 80 |  | 40 |  | 80.4 | 40.3 |  | 0.4 | 0.3
REMARK: THE DIGITAL THERMO HYGROMETER IS CALIBRATED SUCCESSFULLY WITHIN THE SPECIFIED LIMITS
REFERENCE STANDARDS USED
INSTRUMENT |  | Temperature & Humidity Controller |  |  |  | REPORT NO |  | 2021/01/0059
(Traceable To NABL)
MAKE |  | Ambetronics/ (TRH-307) |  |  |  | DUE DATE |  | 2022-01-06
INST ID. NO |  | PPE/C/TEMP/THC/01 |  |  |  | ACCURACY |  | ± 0.3 % F.S / ±1.8% RH
RANGE |  | (-40) to 120 °C / 0-100%RH |  |  |  | LEAST COUNT |  | 0.1°C / 0.1% RH
CALIBRATED BY |  |  |  | CHECKED & APPROVED BY |  |  | REVIEWED BY
NOTE
This report refers only to the particular item calibrated at site/laboratary.
Pin Point Engineers is not liable for any change in calibration data & performance specifications on account of malfunctioning the standards/instrument/equipment covered by this report due to damage caused to it after issuance of this report.
The Standard(s) used for calibration is/are calibrated by using reference standards traceable to National/International standards.
The calibration results reported are valid at the time of and under the conditions of measurement.
This report may not be reproduced, except in full, without the prior permission of "Pin Point Engineers".
Any correction in this certificate invalidated the certificate.
CALIBRATION CERTIFICATE NO: PPE078DTHM1121/036
CLIENT NAME: |  |  | Rusan Pharma Limited
ADDRESS: |  |  | Khasra No.: 122 MI, Central Hope Town, Behind Pharma City, Selaqui,
Dehradun-248197, INDIA
ENVIRONMENTAL CONDITION: |  |  | AMBIENT TEMPERATURE & HUMIDITY
INSTRUMENT DETAILS
INSTRUMENT |  |  | Digital Thermo Hygrometer |  |  | DEPARTMENT |  | QC
MAKE |  |  | HTC |  |  | LOCATION |  | Observation Room
RANGE |  |  | -10 to 50°C/ 10 to 99% RH |  |  |  |  | ( Microbiology)
ACCURACY |  |  | ± 1°C / ± 5% RH |  |  | TAG NO |  | DTHM-106
LEAST COUNT |  |  | 0.1°C & 1% RH |  |  | CAL. DATE |  | 2021-11-02
UNIT OF MEASUREMENT |  |  | %RH & °C |  |  | DUE ON |  | 2022-11-01
CALIBRATION PROCEDURE NUMBER | [5 empty] | PPE/SOP/TEMP/DTH/01-00
OBSERVATION TABLE
Sr. No | STANDARD READINGS ON UNIT UNDER CALIBRATION |  |  |  | OBSERVED READING ON MASTER INSTRUMENT |  |  | ERROR (±)
| RH |  | TEMP. |  | RH | TEMP. |  | RH | TEMP.
1 | 40 |  | 10 |  | 40.1 | 10.2 |  | 0.1 | 0.2
2 | 60 |  | 30 |  | 60.3 | 30.2 |  | 0.3 | 0.2
3 | 80 |  | 40 |  | 80.4 | 40.3 |  | 0.4 | 0.3
REMARK: THE DIGITAL THERMO HYGROMETER IS CALIBRATED SUCCESSFULLY WITHIN THE SPECIFIED LIMITS
REFERENCE STANDARDS USED
INSTRUMENT |  | Temperature & Humidity Controller |  |  |  | REPORT NO |  | 2021/01/0059
(Traceable To NABL)
MAKE |  | Ambetronics/ (TRH-307) |  |  |  | DUE DATE |  | 2022-01-06
INST ID. NO |  | PPE/C/TEMP/THC/01 |  |  |  | ACCURACY |  | ± 0.3 % F.S / ±1.8% RH
RANGE |  | (-40) to 120 °C / 0-100%RH |  |  |  | LEAST COUNT |  | 0.1°C / 0.1% RH
CALIBRATED BY |  |  |  | CHECKED & APPROVED BY |  |  | REVIEWED BY
NOTE
This report refers only to the particular item calibrated at site/laboratary.
Pin Point Engineers is not liable for any change in calibration data & performance specifications on account of malfunctioning the standards/instrument/equipment covered by this report due to damage caused to it after issuance of this report.
The Standard(s) used for calibration is/are calibrated by using reference standards traceable to National/International standards.
The calibration results reported are valid at the time of and under the conditions of measurement.
This report may not be reproduced, except in full, without the prior permission of "Pin Point Engineers".
Any correction in this certificate invalidated the certificate.
CALIBRATION CERTIFICATE NO: PPE078DTHM1121/037
CLIENT NAME: |  |  | Rusan Pharma Limited
ADDRESS: |  |  | Khasra No.: 122 MI, Central Hope Town, Behind Pharma City, Selaqui,
Dehradun-248197, INDIA
ENVIRONMENTAL CONDITION: |  |  | AMBIENT TEMPERATURE & HUMIDITY
INSTRUMENT DETAILS
INSTRUMENT |  |  | Digital Thermo Hygrometer |  |  | DEPARTMENT |  | QC
MAKE |  |  | HTC |  |  | LOCATION |  | Instrument Lab-I
RANGE |  |  | -10 to 50°C/ 10 to 99% RH |  |  |  |  | Primary WS
ACCURACY |  |  | ± 1°C / ± 5% RH |  |  | TAG NO |  | DTHM-068
LEAST COUNT |  |  | 0.1°C & 1% RH |  |  | CAL. DATE |  | 2021-11-02
UNIT OF MEASUREMENT |  |  | %RH & °C |  |  | DUE ON |  | 2022-11-01
CALIBRATION PROCEDURE NUMBER | [5 empty] | PPE/SOP/TEMP/DTH/01-00
OBSERVATION TABLE
Sr. No | STANDARD READINGS ON UNIT UNDER CALIBRATION |  |  |  | OBSERVED READING ON MASTER INSTRUMENT |  |  | ERROR (±)
| RH |  | TEMP. |  | RH | TEMP. |  | RH | TEMP.
1 | 40 |  | 10 |  | 40.1 | 10.1 |  | 0.1 | 0.1
2 | 60 |  | 30 |  | 60.2 | 30.1 |  | 0.2 | 0.1
3 | 80 |  | 40 |  | 80.3 | 40.3 |  | 0.3 | 0.3
REMARK: THE DIGITAL THERMO HYGROMETER IS CALIBRATED SUCCESSFULLY WITHIN THE SPECIFIED LIMITS
REFERENCE STANDARDS USED
INSTRUMENT |  | Temperature & Humidity Controller |  |  |  | REPORT NO |  | 2021/01/0059
(Traceable To NABL)
MAKE |  | Ambetronics/ (TRH-307) |  |  |  | DUE DATE |  | 2022-01-06
INST ID. NO |  | PPE/C/TEMP/THC/01 |  |  |  | ACCURACY |  | ± 0.3 % F.S / ±1.8% RH
RANGE |  | (-40) to 120 °C / 0-100%RH |  |  |  | LEAST COUNT |  | 0.1°C / 0.1% RH
CALIBRATED BY |  |  |  | CHECKED & APPROVED BY |  |  | REVIEWED BY
NOTE
This report refers only to the particular item calibrated at site/laboratary.
Pin Point Engineers is not liable for any change in calibration data & performance specifications on account of malfunctioning the standards/instrument/equipment covered by this report due to damage caused to it after issuance of this report.
The Standard(s) used for calibration is/are calibrated by using reference standards traceable to National/International standards.
The calibration results reported are valid at the time of and under the conditions of measurement.
This report may not be reproduced, except in full, without the prior permission of "Pin Point Engineers".
Any correction in this certificate invalidated the certificate.
CALIBRATION CERTIFICATE NO: PPE078DTHM1121/038
CLIENT NAME: |  |  | Rusan Pharma Limited
ADDRESS: |  |  | Khasra No.: 122 MI, Central Hope Town, Behind Pharma City, Selaqui,
Dehradun-248197, INDIA
ENVIRONMENTAL CONDITION: |  |  | AMBIENT TEMPERATURE & HUMIDITY
INSTRUMENT DETAILS
INSTRUMENT |  |  | Digital Thermo Hygrometer |  |  | DEPARTMENT |  | QC
MAKE |  |  | HTC |  |  | LOCATION |  | Media Fill Incubation
RANGE |  |  | -10 to 50°C/ 10 to 99% RH
ACCURACY |  |  | ± 1°C / ± 5% RH |  |  | TAG NO |  | DTHM-072
LEAST COUNT |  |  | 0.1°C & 1% RH |  |  | CAL. DATE |  | 2021-11-02
UNIT OF MEASUREMENT |  |  | %RH & °C |  |  | DUE ON |  | 2022-11-01
CALIBRATION PROCEDURE NUMBER | [5 empty] | PPE/SOP/TEMP/DTH/01-00
OBSERVATION TABLE
Sr. No | STANDARD READINGS ON UNIT UNDER CALIBRATION |  |  |  | OBSERVED READING ON MASTER INSTRUMENT |  |  | ERROR (±)
| RH |  | TEMP. |  | RH | TEMP. |  | RH | TEMP.
1 | 40 |  | 10 |  | 40.2 | 10.3 |  | 0.2 | 0.3
2 | 60 |  | 30 |  | 60.2 | 30.3 |  | 0.2 | 0.3
3 | 80 |  | 40 |  | 80.4 | 40.3 |  | 0.4 | 0.3
REMARK: THE DIGITAL THERMO HYGROMETER IS CALIBRATED SUCCESSFULLY WITHIN THE SPECIFIED LIMITS
REFERENCE STANDARDS USED
INSTRUMENT |  | Temperature & Humidity Controller |  |  |  | REPORT NO |  | 2021/01/0059
(Traceable To NABL)
MAKE |  | Ambetronics/ (TRH-307) |  |  |  | DUE DATE |  | 2022-01-06
INST ID. NO |  | PPE/C/TEMP/THC/01 |  |  |  | ACCURACY |  | ± 0.3 % F.S / ±1.8% RH
RANGE |  | (-40) to 120 °C / 0-100%RH |  |  |  | LEAST COUNT |  | 0.1°C / 0.1% RH
CALIBRATED BY |  |  |  | CHECKED & APPROVED BY |  |  | REVIEWED BY
NOTE
This report refers only to the particular item calibrated at site/laboratary.
Pin Point Engineers is not liable for any change in calibration data & performance specifications on account of malfunctioning the standards/instrument/equipment covered by this report due to damage caused to it after issuance of this report.
The Standard(s) used for calibration is/are calibrated by using reference standards traceable to National/International standards.
The calibration results reported are valid at the time of and under the conditions of measurement.
This report may not be reproduced, except in full, without the prior permission of "Pin Point Engineers".
Any correction in this certificate invalidated the certificate.
CALIBRATION CERTIFICATE NO: PPE078DTHM1121/039
CLIENT NAME: |  |  | Rusan Pharma Limited
ADDRESS: |  |  | Khasra No.: 122 MI, Central Hope Town, Behind Pharma City, Selaqui,
Dehradun-248197, INDIA
ENVIRONMENTAL CONDITION: |  |  | AMBIENT TEMPERATURE & HUMIDITY
INSTRUMENT DETAILS
INSTRUMENT |  |  | Digital Thermo Hygrometer |  |  | DEPARTMENT |  | QC
MAKE |  |  | Instrumentics
RANGE |  |  | -10 to 50°C/ 10 to 99% RH |  |  | LOCATION |  | Control Sample Room
ACCURACY |  |  | ± 1°C / ± 5% RH |  |  | TAG NO |  | DTHM-003
LEAST COUNT |  |  | 0.1°C & 1% RH |  |  | CAL. DATE |  | 2021-11-11
UNIT OF MEASUREMENT |  |  | %RH & °C |  |  | DUE ON |  | 2022-11-10
CALIBRATION PROCEDURE NUMBER | [5 empty] | PPE/SOP/TEMP/DTH/01-00
OBSERVATION TABLE
Sr. No | STANDARD READINGS ON UNIT UNDER CALIBRATION |  |  |  | OBSERVED READING ON MASTER INSTRUMENT |  |  | ERROR(±)
| RH |  | TEMP. |  | RH | TEMP. |  | RH | TEMP.
1 | 40 |  | 10 |  | 40.1 | 10.2 |  | 0.1 | 0.2
2 | 60 |  | 30 |  | 60.4 | 30.2 |  | 0.4 | 0.2
3 | 80 |  | 40 |  | 80.3 | 40.2 |  | 0.3 | 0.2
REMARK: THE DIGITAL THERMO HYGROMETER IS CALIBRATED SUCCESSFULLY WITHIN THE SPECIFIED LIMITS
REFERENCE STANDARDS USED
INSTRUMENT |  | Temperature & Humidity Controller |  |  |  | REPORT NO |  | 2021/01/0059
MAKE |  | Ambetronics/ (TRH-307) | [5 empty] | (Traceable to NABL)
INST ID. NO |  | PPE/C/TEMP/THC/01 |  |  |  | DUE DATE |  | 2022-01-06
RANGE |  | (-40) to 120 °C / 0-100%RH |  |  |  | ACCURACY |  | ± 0.3 % F.S / ±1.8% RH
[6 empty] | LEAST COUNT |  | 0.1°C / 0.1% RH
CALIBRATED BY |  |  |  | CHECKED & APPROVED BY |  |  | REVIEWED BY
NOTE
This report refers only to the particular item calibrated at site/laboratary.
Pin Point Engineers is not liable for any change in calibration data & performance specifications on account of malfunctioning the standards/instrument/equipment covered by this report due to damage caused to it after issuance of this report.
The Standard(s) used for calibration is/are calibrated by using reference standards traceable to National/International standards.
The calibration results reported are valid at the time of and under the conditions of measurement.
This report may not be reproduced, except in full, without the prior permission of "Pin Point Engineers".
Any correction in this certificate invalidated the certificate.